## Slide 1: Per- and Polyfluoroalkyl Substances (PFAS) Contamination in Materials Used to Collect, Prepare, and Analyze for PFAS in the Environment
Danielle Kleinmaier, Lawrence Zintek
U.S. Environmental Protection Agency
Region 5 Laboratory

## Slide 2: Disclaimer
Reference herein to any specific commercial product, process, or service by trade name, trademark, manufacturer, or otherwise does not constitute or imply its endorsement, recommendation, or favoring by the United States government.
The views and opinions of the authors expressed herein do not necessarily state or reflect those of the United States government or the United States Environmental Protection Agency and shall not be used for advertising or product endorsement purposes.

## Slide 3: Background
PFAS are commonly observed, persistent, and appear to be resistant to many treatment processes (“forever chemicals”)
Compounds manufactured to make products more resistant to stains, grease, and water
Found in many products:
  Foams used for fire suppression
  Non-stick cookware
  Waterproof and stain-resistant textiles
  Water and oil resistant papers
  Metal plating and etching fluids

## Slide 4: Historically Common Contamination Sources in the Field or Lab
PTFE- or FEP-containing products (e.g. Teflon®)
PVDF- or ETFE-containing products (e.g. certain kinds of tubing)
Waterproof, water-resistant, or stain-resistant clothing/products (e.g. outdoor wear, lab coats, fast food wrappers)
Certain personal care products (e.g. cosmetics, lotions)
Certain insect repellants and sunscreens
Plastic clipboards, binders, or hardcover spiral books
Post-it® notes
Recycled paper products (e.g. paper towels, notebook paper)
Chemical (blue) ice packs
Disposable glass pipettes
Aluminum foil
Kim® wipes
Latex gloves

## Slide 5: ASTM Standards D7979 and D7968
Developed at Region 5 lab for non-drinking water and solid matrices
  ~3,000 field samples analyzed since 2012 (+ ~2,800 QC samples)
Direct injection methods (i.e. “dilute and shoot”)
  Minimal sample manipulation reduces prep time, minimizes risk of blank contamination
External standard methods
  Labelled surrogates for most target analytes, used strictly to evaluate method performance
Analysis by LC/MS/MS
  Confirmatory transitions for most target analytes
  Ion ratios calculated to support qualitative IDs

## Slide 6: Current Region 5 Lab PFAS Analyte List
Target Analyte | Reporting Limit in Water (ng/L) | Reporting Limit In Soil (ng/kg) | Labelled Surrogate
PFBA, PFPeA | 50 | 125 | x
PFBS | 10 | 25 | x
PFPeS | 10 | 25 | 
PFHxA | 10 | 50 | x
PFHxS, PFHpA, PFOA, PFNA, PFDA, PFUnA, PFDoA, PFTreA | 10 | 25 | x
PFHpS, PFNS, PFDS | 10 | 25 | 
PFOS | 10 | 30 | x
PFTriA | 10 | 25 | 
FOSA, 4:2 FTS, 6:2 FTS, 8:2 FTS, N-EtFOSAA, N-MeFOSAA | 10 | 25 | x

## Slide 7: ASTM Standards – Sample Preparation
D7979 (PFAS in Water Method)
5 mL sample (no sub-sampling!)
Spike with surrogates, dilute 1:1 with 5 mL methanol, shake for 2 minutes
Filter through pre-rinsed hydrophilic polypropylene membrane, acidify with acetic acid (pH 3 – 4)
10 mL final volume
D7968 (PFAS in Soil Method)
2 gram sub-sample
Spike with surrogates, extract via tumbling with 10 mL 1:1 methanol/water at pH 9 – 10 (adjusted with ammonium hydroxide) for 1 hour
Centrifuge extract, filter through pre-rinsed hydrophilic polypropylene membrane, acidify with acetic acid (pH 3 – 4)
10 mL final volume

## Slide 8: ASTM Standards – Analysis
Parts-per-trillion (ppt) analyte calibration range
21-minute run time
Ternary LC gradient

## Slide 9: Lab Consumables
Lab consumable = lab supply that is considered one-time use, disposable, and inexpensive
Lab consumables in Region 5 lab PFAS methods:
  Polypropylene sampling containers (15 mL and 50 mL conical centrifuge tubes)
    Used also for reagent storage and standard preparation/storage
  2 mL amber glass instrument vials (i.e. autosampler vials)
  Polyethylene screw caps for 2 mL instrument vials
  Polypropylene pipette tips (200 µL, 1 mL, 10 mL)
  Polypropylene filtration disks (with 0.2 µm GHP membrane)

## Slide 10: Erratic Lab Consumables Contamination
In July 2020, erratic PFAS contamination was discovered while processing a data set
  Contamination varied in frequency, concentration, and affected PFAS target analytes
  Such contamination caused the data set to be of unknown quality
  All PFAS preparation/analysis ceased in the lab until the root cause of the contamination could be determined and corrected
  After 4 months of investigation, the following lab consumables used for PFAS were found to erratically contain PFAS:
    2 mL amber glass instrument vials (i.e. autosampler vials)
    15 mL and 50 mL polypropylene conical centrifuge tubes (used for field sampling, reagent storage, and standard preparation/storage)
    10 mL polypropylene pipette tips

## Slide 11: Erratic Lab Consumables Contamination
Several brands, part numbers, and lots of affected consumables were extensively tested
  Other brands and part numbers contained erratic PFAS contamination as well
  Determined in-house self-certification of lots not feasible with limited staff/resources
Short-term solution: all PFAS field samples required to be collected, prepared, analyzed, and reported in triplicate
  Allows for easy identification of false positives or biased-high results
    “T” flag created to notify data users of potential false positives/biased-high results in data reports
  Provides needed assurance to data users that the reported PFAS results are both precise and accurate

## Slide 12: PFAS Contamination Data
98 aqueous field samples (236 in triplicate) have been prepared/analyzed since triplicate requirement enacted in November 2020
Method blanks, reporting limit checks, blank spikes, and triplicate field data all monitored for erratic contamination
  2 method blanks, 2 reporting limit checks, and 2 blank spikes required for each batch of 20 (including triplicates) field samples
  7,776 data points collected and evaluated
Contamination is determined using percent difference between the 3 replicates
  ± 30% difference is limit used to determine if result is a false positive or biased high

## Slide 13: PFAS Contamination Data
Example of False Positive Result
PFOS in field sample 2102012-02 was flagged “T” in the data report to notify the client that the data point is likely a false positive
 |  | ≥ Reported RL |  | > ± 30% Difference (D) Limit |  | 
Sample ID | Analyte | Reported Result (ppt) | Reported RL (ppt) | %D -02 vs. -02RE1 | %D -02 vs. -02RE2 | %D -02RE1 vs. -02RE2
2102012-02 | PFOS | 28.9 | 11.1 | 826 | 820 | -0.64
2102012-02RE1 | PFOS | 3.12 | 10.3 |  |  | 
2102012-02RE2 | PFOS | 3.14 | 10.1 |  |  | 

## Slide 14: PFAS Contamination Data
Chart: Frequency of PFAS Contamination
| PFHpA | PFHxA | PFTriA | PFTreA | PFUnA | PFDoA | 6:2 FTS | PFOS | PFPeA | PFDA | PFBA | PFOA | PFNA | PFHxS | 8:2 FTS | PFBS | PFDS | PFNS | PFHpS | PFPeS | 4:2 FTS
 | 15 | 14 | 12 | 11 | 10 | 9 | 9 | 9 | 7 | 5 | 5 | 5 | 5 | 4 | 2 | 2 | 1 | 1 | 1 | 1 | 1

## Slide 15: PFAS Contamination Data
Chart: Average PFAS Contamination Concentrations
| PFBS | PFPeA | PFDoA | PFHxA | PFBA | PFTreA | 8:2 FTS | 6:2 FTS | PFOA | PFTriA | PFHxS | PFUnA | PFOS | PFHpA | PFNA | PFDA | PFPeS | PFHpS | 4:2 FTS | PFDS | PFNS
 | 500.5 | 200.4714 | 186.0078 | 138.1057 | 121.3 | 87.7582 | 52.9 | 48.0333 | 42.14 | 42.0908 | 27 | 23.444 | 21.57 | 21.1687 | 13.392 | 8.376 | 6.8 | 6.7 | 6.2 | 5.4 | 5.3
129 contaminated data points

## Slide 16: PFAS Contamination Data
Chart: Highest PFAS Contamination Data Points
| PFDoA | PFHxA | PFPeA | PFBS | PFTreA | PFTriA | PFBS | PFTreA | PFBA | PFPeA | PFOA | 6:2 FTS | PFHxA | PFTreA | PFBA | PFPeA | 8:2 FTS | PFBA | PFBA | PFPeA
 | 1550 | 1370 | 875 | 781 | 398 | 236 | 220 | 219 | 214 | 200 | 154 | 148 | 131 | 128 | 127 | 111 | 100 | 98.5 | 89 | 85.4
20 highest contaminated data points

## Slide 17: PFAS Contamination Data – Summary
Additional statistics:
  88% of Region 5 lab PFAS method target analytes affected by continued erratic contamination
  67% of contaminated data points at or near method reporting limits
  On average, 2% of data points from each data set impacted by contamination
Caution should still be taken by data users when evaluating data variability in low-level results in terms of project decision making as replicate PFAS samples are co-located
Long-term solution: consumable vendors provide certified, trace-level (low ppt) PFAS-free products

## Slide 18: Considerations When Requesting PFAS Analysis
 | EPA Method 533 | EPA Method 537.1 | ASTM Standard D7979 | EPA Method 8327
Lowest Target Analyte Calibration Range: | 0.5 - 25 ppb (on-column) | 0.5 - 25 ppb* (on-column) | 5 - 200 ppt (on-column) | 5 - 200 ppt (on-column)
Lowest Demonstrated Reporting Limit: | 10 ppt | 16 ppt | 10 ppt | 10 ppt
Initial Sample Volume: | 250 mL | 250 mL | 5 mL | 5 mL
Final Sample Volume (for Analysis): | 1 mL | 1 mL | 10 mL | 10 mL
*not published, assumed
What analytical method does the lab run?

## Slide 19: Considerations When Requesting PFAS Analysis
What analytical method does the lab run?
  If the method calibrates PFAS analytes in the ppt-range (on-column), it is strongly recommended to request triplicate sampling, preparation, analysis, and reporting for each sampling location
    E.g. ASTM Standards D7979 and EPA Method 8327
  If the method calibrates PFAS analytes in the ppb-range (on-column), the presented data suggests that the currently observed contamination in lab consumables may not impact these methods
    E.g. EPA Methods 533 and 537.1 (providing the analyte calibration range or method RLs are not modified to be lower)
    Also note that consumables used for these methods may differ from those used in D7979 or 8327

## Slide 20: Considerations When Requesting PFAS Analysis
Why not just always use EPA Methods 533 or 537.1?
  These are drinking water methods that were only validated by the EPA to be used for drinking water matrices
  ASTM Standard D7979 and EPA Method 8327 were validated to be used for several aqueous matrices including surface water, groundwater, and wastewater (influent and effluent)
    These direct injection methods also have significantly higher sample throughput rates, even when preparing samples in triplicate, than the labor-intensive drinking water methods
Once consumable vendors begin providing certified, trace-level (low ppt) PFAS-free products, triplicate analysis will no longer be necessary

## Slide 21: Contact Information
Larry Zintek
zintek.lawrence@epa.gov
(312) 886-2925
Danielle Kleinmaier
kleinmaier.danielle@epa.gov
(312) 353-9771